## Summary
Consultation response for
098 Medical devices in plastic or silicone
Criteria generation 3
Consultation period: from 3 December 2025 to 6 February 2026
The Consultation
The proposed requirement for the 098 Medical devices in plastic or silicone was under consultation from 3 December 2025 to 6 February 2026. We received a total of 13 official consultation responses.
We have received responses from 2 Danish, 5 Swedish, 3 nordic countries together and 3 "other" respondents.
All consultation responses are collected in this Excel file originally generated from Forms and received by email.
Most responses were regarding: - What can carry the Nordic Swan Ecolabel? - O5 Silicone - O9-O13 Chemical requirements - O15-O16 Packaging requirements
How to read the Consultation Response document:
Nordic Ecolabelling's responses to the input are collected in row 2 in the sheet "responses and feedback", divided into topics and requirements.
It is possible to sort within the document to make it more readable (e.g. by country). You can also adjust the size of columns or rows if it is difficult to read all the text.
Key Changes in the criteria after consultation:
The title of the criteria has change from “Medical devices in plastic or silicone” to “Medical devices and medicinal products in plastic and silicone”.
O5 Silicone: Added that for small parts that weight max. 2 grams and that are not introduced into the patient during treatment then limit is 1000 ppm for D4, D5 and D6 and this must be declared by the manufacturer/supplier of the silicone. Added alternative to medical-grade silicone, that silicone material can be tested according to ISO 10993 or USP class VI.
O5 Silicone and O14 Silicone in adhesive and surface treatment: Added test method for testing level of D4, D5 and D6.
O8 Silicone production: Emission of CO2 has been replaced by a requirement for information about energy consumption.
015 Packaging materials: Added exception for paper part of protective blister in primary packaging regarding must be FSC/PEFC certification or recycled.
O16 Packaging design for recycling: Requirement for pictograms for recycling and text regarding separation of materials before waste sorting have been deleted.
O17 Safety: Added requirement that the manufacturing site of the final product must be certified with ISO 13485 (Medical Devices - Quality management systems – Requirements for regulatory purposes).
New requirement for energy mapping and energy action plan for manufacturing sites that preform sterilization of the final product.
Thanks to all responding organizations for valuable feedback during this consultation!

## Responding organizations
Organisation | Country
Olympus Sverige AB | Sweden
BD - Becton Dickinson | Switzerland
Baxter | Sweden
Mölnlycke Health Care | Represent all nordic countries
Hollister | Sweden
Fresenius Medical Care AG | Germany
Region Stocholm_Nordic healthcare actors (NIELS working group) | Sweden, Denmark, Norway, Finland
Coloplast A/S | Denmark
Institutet för Miljömedicin, Karolinska Institutet | Sweden
Miljøstyrelsen | Denmark
Wellspect | Sweden
Swedish Medtech, Sailab, Melanor, Medicoindustrien | Sweden, Norway, Finland, Denmark
VinylPlus Healthcare, PVC Forum Sverige and PVC-Informationsrådet | Europe, Sweden, Denmark

## Responses and feedback
Responses | Organisation | Country | General feedback | The criteria document as a whole | What can carry the Nordic Swan Ecolabel? | Requirement O1 Description of the product and manufacturing process | Requirement O2 Halogenated plastics | Requirement O3 Halogenated butyl rubber | Requirement O4 Natural rubber latex | Requirement O5 Silicone | Silicone production - Requirement O6 Emissions of dust and chlorides to air | Silicone production - Requirement O7 Emissions of copper and of zinc to water | Silicone production - Requirement O8 Emissions of CO2 | Requirement O9 Classification of ingoing substances in additives | Requirement O10 Excluded substances in additives | Requirement O11 Classification of adhesive and surface treatment | Requirement O12 Classification of ingoing substances in adhesive and surface treatment | Requirement O13 Excluded substances in adhesive and surface treatment | Requirement O14 Silicone in adhesive and surface treatment | Requirement O15 Packaging materials | Requirement O16 Packaging design for recycling | Requirement O17 Safety | Requirement O18 Customer complaints | Requirement O19 Traceability | Appendices
Consultation response from Nordic Ecolabelling |  |  | The Nordic Swan Ecolabel is the official ecolabel of the Nordic countries supported by all Nordic governments and is the most recognised environmental label in the region. The Nordic Swan Ecolabel is one of the world's toughest environmental certifications, and from our knowledge at this moment, the only ecolabel within this segment. We are aware that many companies in this industry act on a global market and we understand that the Nordic region may have some difficulties influencing the industry on a global scale. However, the vision from the Nordic Council of Ministers is to make "the Nordic region the most sustainable and integrated region in the world" which is well incorporated in our ways of working. This means that we strive for the Nordic region to take the lead in this industry by setting ambitious requirements. The Nordic Swan Ecolabel is only awarded to products and services meeting ambitious environmental requirements. Since Nordic Ecolabelling set strict requirements, only the best products in the market will be able to obtain a certification. In the criteria development process, an important issue is therefore to only include product types and requirements that help distinguish the better products in the market. We need to find features and properties within each specific product type where we can set requirements to distinguish the environmentally better products in each segment. The same goes for specific requirements. We cannot set a specific requirement if that requirement does not contribute to distinguish the environmentally better products. In order for us to distinguish the better products in the market, we need to request more information than legislation usually does. We strive to moving industries by requesting more information and we understand that this could be a long-term task in many cases, but this work will result in better products for both environment and health. Of course, we would like to see an increase of Nordic Swan Ecolabelled products in medical technology, but not at the expense of decreased environmental and health performance. All the documentation are handled as confidential and the sub-suppliers are more than welcome to submit documentation confidentially directly to the Ecolabelling organization. | Thank you for your feedback. We have decided to prolong the current generation 2 of the criteria until 31 August 2028 giving a transition periode of a little more than 2 years to adjust and collect documentation according to generation 3. Answer to Olympus Sverige: The Nordic Swan Ecolabel is a voluntary ecolabel where it is not possible to set an evaluation criteria, but must set mandatory requirements. During the development of the draft criteria there is dialogue with licensees and during the consultation it is the aim to get responses from the industry, NGOs, governments ect. This is important inputs to be able to find the right level of the requirements. Answer to BD: Only products made of min. 90 wt% plastic or silicone (or mix) and that are listen under section "What can carry the Nordic Swan Ecolabel?" can be labelled under these criteria. However, products made of min. 90 wt% plastic/silicone that are not listed, may be added if Nordic Ecolabelling decide to do so after suggestion from interested part. It is not possible to make different levels of the Nordic Swan Ecolabel. See above regarding transition periode. MECO scheme is qualitative because of the diverse products types covered by the criteria. Answer to Baxter: The Nordic Swan Ecolabel is voluntary. See above regarding transition period. Answer to Mölnlycke Health Care: The product types covered by the criteria are selected because the RPSs (see section "Environmental impact of medical devices in plastic or silicone") show that it is relevant and possible to set requirements for the materials, such as alternative materials to PVC or production of silicone. The RPSs also show other relevant and possible areas of requirements such as chemical substances and packaging. The Nordic Swan Ecolabel is a voluntary ecolabel, which set requirements that go beyond legislation, which of course must always be complied with. The Nordic Swan Ecolabel is always product specific, a lot of licensees have both ecolabelled and non-ecolabelled products in their portfolio. Answer to Fresenius Medical Care: The Nordic Council of Ministers has determined the overall rules for the Nordic Ecolabelling and have stated that the criteria needs to be evaluated regularly. The consultation is a very important part of the revision and for the level of requirements. Swedish Medtech, Sailab, Melanor, Medicoindustrien: The Nordic Swan Ecolabel is a voluntary ecolabel, which set requirements that go beyond legislation. Legislation must of course always be complied with. Because it is not mandatory for products to be Nordic Swan Ecolabelled, you can choose not to do so if you find it too administratively burdensome to apply for and maintain your license. Nordic Ecolabelling's criteria are aimed at, but not limited to the Nordic countries. Therefore, once you have produced a good medical device with the Nordic Swan Ecolabel that meets high sustainability requirements, it is obvious to sell it not only to the Nordic region but also to the rest of the world. We are very aware that Nordic Ecolabelling's requirements in these criteria do not lead to undesirable consequences for the product's functionality and safety, which is why the requirements have also been developed in close dialogue with manufacturers of medical devices. The development of criteria in Nordic Ecolabelling is always product specific, the requirements are adapted to the product types covered by the criteria. Dialogue with the industry and public consultation are important parts of the criteria development. The criteria only cover the product types described in the section "What can carry the Nordic Swan Ecolabel?". These product types are the types specific listed under the categories masks, flexible tubes, bags, plugs and syringes. These product types are generally class I or II under MDR. The product types covered by the criteria are selected because the RPSs (see section "Environmental impact of medical devices in plastic or silicone") show that it is relevant and possible to set requirements for the materials, such as alternative materials to PVC or production of silicone. The RPSs also show other relevant and possible areas of requirements such as chemical substances and packaging. Silicone is allowed, but there are requirements to the production (for products made of min. 80 wt% silicone), to additives and to the level of D4, D5 and D6. VinylPlus Healthcare, PVC Forum Sverige and PVC-Informationsrådet: Please see answer form Nordic Ecolabelling under Requirement O2. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation, but have changed the name of the criteria from “Medical devices in plastic or silicone” to “Plastic or silicone products in medical technology”. Answer to BD: The assessment from Nordic Ecolabelling is that products included in the IVD regulation are out of scope of these criteria. The IVD regulation includes for example diagnostic test kits, laboratory instruments and diagnostic softwares. Products that not fulfill the definiton that minimum 90% of the product is made of plastic or silicone and is used to storage, transfer or transport of fluid, gas or medicine; or for preventing leakage of fluids from the body. If there are any products included in IVD regulation which fulfill the part of the product group definition described above and it shows that we can set requirements to distinguish the environmentally better products on the market, Nordic Ecolabelling might consider adding the specific product and the regulation to the criteria document upon request. For now, it is only the listed products that is possible to Nordic Swan Ecolabel. Products that fulfill the product group definition and where our requirements can make a difference within the product category can be added to the product group upon request from the industry. Answer to Fresenius Medical Care: We have decided to change the title of the criteria to the suggested one. Answer to NIELS working group: The silicone requirements are only applied for products with 80% or more silicone. A product can e.g. be 50% plastic and 50% silicone and thus have met the product group definition, but the silicone production requirements are not applicable. Laryngeal mask is added to the product group definition. Bottle teats are normally not included in Regulation (EU) 2017/745 on medical devices, therefore they are not included in the criteria. The investigation of hazardous waste containers, needle buckets, dosing box, medicine cup, washset bowl, sampling tubes and bowls/bottles/jars/tubes for laboratory culture/tests showed that there is a low RPS. This means that we haven't found features and properties within this product category where we can set requirements to distinguish the environmentally better products and therefore they are not included in the criteria. Plugs were not investigated in the revision, because they were not suggested as a new product type when the project group requested suggestions for new product types. However, if we receive a request from a potential license holder, then it would be possible for Nordic Ecolabelling to investigate if can be added to the product group. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to Olympus Sverige: All information and documentation that is submitted to Nordic Ecolabelling is handled confidential, besides our public consultations. However, it is always possible to submit additional information that will be held confidential, in conjunction with the public consultation. It is also always possible to sign a NDA. Answer to BD: Nordic Ecolabelling strives to set mostly product specific requirements. While both ISO 14001 and ISO 50001 are general standards set on company level, we have not added them to the product specific requirements. An inspection visit is always mandatory for each manufacturing site. All the fees are available at our website. Answer to NIELS working group: Nordic Ecolabelling require safety data sheets for all additives, adhesives and surface treatment according to current legislation. The person filling and signing the appendix should have sufficient knowledge to fill in the appendix and is the responsible person representing their organization. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to BD: Halogenated plastics are not allowed. There is no lower threshold. Answer to VinylPlus et al: Nordic Ecolabelling has updated the background documentation text and references for the argumentation for not accepting PVC as a material in medical devices. The requirement itself has not been updated. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to NIELS working group: For the product types suggested to be covered in the criteria it is possible to have alternative materials to natural rubber latex. In the case that Nordic Ecolabelling gets a request to expand the citeria to cover a product type consisting of natural rubber latex, then the suggested requirements to natural rubber latex will be considered. | Thank you for your feedback. The requirement has been changed after the consultation. Based on dialog with manufacturers it has been added that small parts (now defined as parts max. 2 gram and that are not introduced into the patient during treatment) the limit for D4, D5 and D6 is 1000 ppm in the silicone material, which must be declared by the manufacturer/supplier of the silicone. For other parts silicone can be tested according to ISO 10993 or USP class VI, as an alternativ to medical-grade silicone. Also, an additional test metode for the content for D4, D5 and D6 has been added. Answer to DB: The limits of 100 ppm/1000 ppm is for the silicone material. Requirments for lubricants are found under requirement O14. Answer to NIELS working group: We have considered linear siloxanes, but have for this generation focused on test for the amount of D4, D5 and D6 (except for small parts as decribed above). Answer to Coloplast: We have added you suggested test metode. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. | Thank you for your feedback. After the consultation the following requirement text has been added regarding concentration limits for copper and zink: "However, if national or local legislation is stricter, then these levels must be met". Answer to Miljøstyrelsen: Please see above. | Thank you for your feedback. The requirement has been changed after the consultation. The requirement for emission of CO2 has been replaced by a requirement for information about energy consumption. Answer to NIELS working group: Please see above. Waste water is covered in req. O7. | Thank you for your feedback. After the consultation it has been deleted that SDS for additives must be send. Answer to BD: After the consultation it has been deleted that SDS for additives must be send. The manufacturer/supplier of the polymer material must completed and signed appendix 5 regarding requirement O9. Answer to NIELS working group: We also include substances that are self-classified at the same level as the harmonised classifications. In addition to classification we have the excluded substances lists, which cover substances that may not be classified but have unwanted properties. Answer to Institutet för Miljömedicin: "Hazardous to the aquatic environment" and "acute toxicity" are included for the requirement for adhesive and surface treatment (O11, the whole raw material), but not in inoging substances in additives in polymer materials (O9) because of the risk for the end user and exposure to the aquatic environment, where additives in polymer materials with these classifications are expected to have a lower risk compared to adhesives and surface treatments. These classification are also not included in requirement O12 because the requirement refers to classification of ingoing substances in the adhesive and surface treatment, whilst requirement O11 refers to the classification of the whole adhesive and surface treatment. This is also the case regarding the "STOT"-classifications. Answer to Miljøstyrelsen: Clarifications in the background document has been made. However, every manufacturer is free to add the additives as they deem necessary for each product. Nordic Ecolabelling cannot set detailed requirements for which additives are relevant in each specific product, our criteria are only focusing on ensuring that the additives fulfill each requirement. | Thank you for your feedback. After the consultation links have been added to “Endocrine Disruptor Lists" List I, II and III. The exemption under halogenated organic compounds has been deleted and under metals zinc has been deleted. Answer to Institutet för Miljömedicin: Links have been added to “Endocrine Disruptor Lists" List I, II and III. Quaternary ammonium compounds and Volatile aromatic compounds have not been included because they are assessed as less relevant for additives in polymer materials. Regarding background text: General: The requirement is for additives in polymer materials which include plastic and silicone. For bisphenols "now" has been deleted. Answer to Miljøstyrelsen: It is correct that PFAS is a group under halogenated organic compounds. We have nevertheless chosen to have PFAS as a separate point to clearly mark that they are prohibited. The exception for pigments has been deleted, because it is not relevant. Answer to Wellspect: All the relevant lists are presented in each specific appendix. It is possible for suppliers to submit confidential documentation directly to Nordic Ecolabelling. It is also always possible to sign a NDA. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to BD: A SDS in accordance with REACH Regulation (EC) No 1907/2006 and an appendix 6 completed and signed by the manufacturer/supplier. Answer to Institutet för Miljömedicin: See the answer in O9. Answer to Miljøstyrelsen: We consider that the risk of not fully hardened acrylates is low due to manufacturers detailed procedures and instructions in their processes. Introducing such a testing requirement would cause major additional testing for the manufacturers with high analytical uncertainty, and therefore we do not see it as justified to introduce such a requirement. Furthermore, if the adhesive is not completely cured it does not perform its intended function, which in itself would cause a quality issue that is taken into account in other parts of the manufacturing process. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to Institutet för Miljömedicin: See the answer in O9. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. Answer to NIELS working group: UV-cured adhesives may be acrylates-based, and there is an increasing number of acrylates that are being classified as environmentally hazardous and skin sensitizing. During the curing process, the acrylates monomers react with each other, so they do not pose an environmental or health hazard in the finished cured adhesive nor in the final product. Therefore, an exemption has been made for UV-cured acrylates-based adhesives, but the exemption only applies if the classified chemical is used in a closed production system, where there is no direct contact/exposure between worker and the chemical product. Using UV-curing have several advantages as it provides durable products and does not require solvents (low VOC-content). Answer to Institutet för Miljömedicin: See the answer in O10. Answer to Wellspect: See the answer in O10. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. | Thank you for your feedback. After the consultation the requirement and background text for board and paper have been updated with more precise text. Definition of recycled material has been added in the section for definition. The paper part of protective blister in primary packaging has been exempt. Answer to BD: It is possible to use FSC- or PEFC-certified packaging that are not recycled. In primary packaging all metal layer for barrier purposes to preserve a sterile barrier, minimize evaporation and/or improve product shelf life are exempted. For other questions please see under O16. Answer to NIELS working group: We have chosen not to include packaging vs. weight of product because the diverse product types covered by the criteria and safety aspects that must be considered for medical devices e.g. if sterility is needed. Filling material have not been considered, but may in next generation of the criteria. Documentation for FSC/PEFC can be e.g. invoices or delivery notes from suppliers. Answer to Coloplast: Please see note in top. Answer to Miljøstyrelsen: Beside halogenated plastics, also oxo-degradable plastic and biodegradable plastic must not be used. The definitions of primary, secondary and tertiary packaging do not follow the definitions from Packaging and Packaging Waste Regulation, because dialogue with the industry showed that these definition was not accurate enough to be used in these criteria for medical devices. | Thank you for your feedback. After the consultation requirement for pictograms and text regarding separation of materials before waste sorting have been deleted. Answer to BD: PFAS means all PFASs acording to the definition in the criteria that are intentionally added or used in surface treatment of the packaging. Carbon black must not added to the plastic materials of the packaging, it do not include ink for printing. Please also see above. Answer to Coloplast: Please see note in top. Answer to Miljøstyrelsen: The applicant must send picture/artwork of packaging and description of how different materials can be separated as documentation. Danish udvidede producentansvarsregler is not very specific regarding seperation of materials, however this is stated in Bilag 2, Væsentlige krav til emballagers sammensætning og mulighederne for genbrug og nyttiggørelse, herunder genanvendelse, jf. § 4, 3. Krav i forbindelse med nyttiggørelse af emballage: "1) Nyttiggørelse af emballage i form af materialegenanvendelse. Emballagen skal fremstilles på en sådan måde, at det er muligt at genanvende en vis vægtprocent af de anvendte materialer til fremstilling af salgbare produkter under hensyn til de gældende standarder i Fællesskabet. Fastsættelsen af denne procentdel kan variere afhængigt af, hvilken type materiale emballagen består af". | Thank you for your feedback. After the consultation requirement for that the manufacturing site of the final product must be certified according to ISO 13485 or EN ISO 13485 “Medical devices – Quality management systems – Requirements for regulatory purposes” has been added. It has been deleted that the notified body of the CE marking must be visible on the label. Answer to NIELS working group: It has been deleted that the notified body of the CE marking must be visible on the label. Answer to Coloplast: Please see note in top. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. | Thank you for your feedback. We have decided to keep the requirement as suggested in the consultation. | Thank you for your feedback Answer to NIELS working group: Requirement O8 for emission of CO2 has been replaced by a requirement for information about energy consumption. A template for information about energy consumption will be available when the criteria are published.
| Olympus Sverige AB | Sweden |  | Overall, this criteria document is problematic. Many companies operating in this market are global actors and are therefore unable to comply with local certification schemes such as the Nordic Swan. While the underlying criteria may be relevant and appropriate, they should primarily be introduced as evaluation criteria in tenders rather than as minimum requirements, at least until the market is sufficiently mature to meet them. A structured and transparent market dialogue is essential when transitioning criteria from evaluation-based to mandatory requirements. This ensures that suppliers are genuinely able to comply with and document fulfilment of the criteria, rather than merely confirming compliance in order to participate in a tender. There is broad consensus across the industry that sustainability is a high priority. However, it is important to acknowledge the significant regulatory burden companies have faced in recent years. For example, the transition from MDD to MDR has required substantial focus and resources to ensure product compliance, safeguard patient safety, and maintain a stable supply to healthcare systems. This regulatory shift has limited the capacity of many companies to simultaneously adapt to additional certification requirements. | Please see overall comment | Some of the requested information under this point is confidential information that cannot be shared with third party. | PVC is still in use in some medical products due to their flexibility, and lack of other good alternatives. Until better alternative is available, this shouldn't be banned.
| BD - Becton Dickinson | Switzerland |  | 1) What additional types of products are in scope? Many medical devices are made with mix of materials, so not exclusively plastics. Would blood sample tubes be in scope? Would IV catheters with needles? Will they fit here if plastic is not greater than 90% of weight? 2) Provide transparency of scorecard and perhaps introduce levels of label, e.g. gold, silver, bronze 3) Also consider company's SBTi validated targets, Supply chain optimization services, other commitments. 3) Some of the input from customers/test reports may be time consuming. What would be the timelines? For the MECO scheme - it is a qualitative method not quantitative? Have you considered the disadvantages? | p4: "The products must be included under the Regulation (EU) 2017/745 on medical devices or EU Medicinal Products Directive (2001/83/EC). The products can either be single-use or reusable" -> why not also include IVDs? Strong recommendation to include products covered under the IVD regulation as well. p5: What is the process to add new products going forward? Please mention this clearly in the document. As it is, only the product types (e.g. “anaesthesia mask”) which are listed under each group (e.g. “masks”) in the product group definition can be labelled. In section 2.1 under Further background: suggest to reference "saline (not salt water) and medication in prefilled syringes" | Manufacturing requirements are missing overall. The only aspect mentioned (but too low RPS) is energy consumption (p58). However, nothing else is mentioned on manufacturing. Strongly recommended to include under 5.2 the requirement for the plant to have ISO14001 and proof of renewable energy usage. Suggestion to also include or at least consider ISO 50001 certification for Energy management. If manufacturing plant audit is no longer included, what type of declarations will replace the audit? We strongly recommend keeping the plant audit, as it is the best way to give the picture of sustainability rigor applied in the manufacturing process. 1) Plant visit is essential to have clarity on standards and processes. For manufacturers outside EU, if considered will assume the cost for Nordic Swan group travel. Provide tiers of cost fees. | Does not allowed mean - Not formulated with or is there a threshold like 0.1% ? |  |  | D4, D5, D6- 100 ppm on what level - Device? Can we keep it to the eu MDR threshold? Method provided is for elastomer. What about the usage as lubricant? |  |  |  | Would you require raw material SDS to be shared? |  | SDS or test reports or both needed for verification? |  |  |  | PFAS: Can we include a threshold of 0.1%w/w? Can we allign with requirements of eu PPWR. Multi-Material Packaging Instruction is provided for customers requesting the data; How would it be checked in this case? Carbon black pigments- Do you mean in ink or is it any other application? Metals-Some medical devices require thicker aluminum foil for protection and barrier properties. Recommend to have 10 um instead of 5 um here For Minimum 70 wt% of the paper- and cardboard in the packaging must be recycled* or the packaging must be FSC- or PEFC-certified.- While this is achievable for grouped packaging, cartons used for transportation packaging (where the strength requirement are critical) may not be able to use more than 40-50% recycled content without negatively affecting the performance off the box during the shipping. | PFAS and others- Can we allign with requirements of eu PPWR. Pictograms- would be in conflict with MDR/IVDR or other regulations compared to using what they recommend with DIN or national symbols.
| Baxter | Sweden |  | We fully support the ambition to strengthen environmental frameworks; however, applying the Nordic Swan Ecolabel requirements to lifesaving medical devices presents significant risks and it conflicts with focus on supply‑chain resilience. Medical devices are already governed by the EU Medical Device Regulation (MDR), one of the world’s most stringent regulatory systems. Introducing a separate Nordic‑only ecolabel framework creates fragmentation, contradicts the goal of regulatory alignment within the EU, and places disproportionate compliance burdens on a very small regional market in a global context. This risks discouraging manufacturers from supplying critical devices to the Nordics, ultimately reducing patient access. Furthermore, the current five‑year update cycle of the Nordic Swan criteria is incompatible with the long product lifecycles, validation requirements, and regulatory constraints inherent to medical devices. Frequent changes jeopardize stability in supply chains and product availability. If we aims to strengthen resilience and ensure secure access to essential medical technologies, environmental harmonization must take place at EU level—not through isolated Nordic labeling. We encourage collaboration toward EU‑wide frameworks that balance environmental progress with patient safety, regulatory coherence, and supply continuity.
| Mölnlycke Health Care | Represent all nordic countries |  | From our perspective and how we produce we are not interested in using the Nordic swan ecolabel, this is due to several reasons and here are some of them: 1. Limited Advancement of Environmental Performance The current criteria for the Nordic Swan Ecolabel largely align with existing EU chemical regulations, which our company and products already comply with. As a result, the label does not substantially drive additional environmental improvements for us. Pursuing certification would primarily create duplicate administrative work and increase costs for our customers, without delivering meaningful environmental benefits. 2. Restriction on Natural Rubber Latex Current ecolabel criteria exclude natural rubber latex due to concerns related to allergies and deforestation. However, latex can be sourced responsibly through robust certification systems such as FSC, which ensure sustainable forestry practices, strong ethical standards, and independent third party audits (e.g., Sedex). The upcoming implementation of the EUDR will further strengthen traceability and oversight. We offer both latex and fossil based alternatives, and our latex products have a significantly lower climate footprint. Enabling the use of responsibly sourced latex is therefore essential as we work to reduce reliance on fossil resources. 3. Inconsistent Application Across Our Product Portfolio Since the label cannot be applied to all of our products, its use would lead to inconsistencies in reporting, communication, and customer understanding. This fragmented applicability also makes it challenging to establish consistent and relevant sustainability requirements across our full product range. 4. Focus on Global, Not Local, Frameworks We operate in a global market, and therefore prioritize internationally recognized certifications and rating systems that more effectively support our sustainability strategy. Frameworks such as FSC, ISCC, SBTi, and Ecovadis provide robust global standards that help ensure sustainable material sourcing and drive continuous improvement across the entire organization. These tools support evidence based progress toward our environmental and net zero goals, which is more aligned with our long term ambitions than adopting a regional ecolabel.
| Hollister | Sweden |  | The criteria is more structured and clearer, including the appendixes. Relevant criteria update in line with industry and product innovations, environmental advancements to alternative materials, while also staying relevant to product composition and technical needs to create high quality, safe medical devices | Clear and thorough definitions and examples. good to include MDR and Medicinal Products Directive. | Clear and thorough requirements | in time with PFAS regulations, it is very timely and relevant to include this requirement in both product and packaging. Also good that the PVC requirement is kept as there are several other alternatives available for the industry to use which has a lower environmental impact | good to include | good to include | Interesting to read the background research done on Silicone and its environmental impact which are the basis of the new requirements. It aligns with the advancement of the criteria. | see above | see above | see above | very clear to read. well aligned with regulations and exemptions seem relevant and well-researched | very clear to read. well aligned with regulations and exemptions seem relevant and well-researched | very clear to read. well aligned with regulations and exemptions seem relevant and well-researched | very clear to read. well aligned with regulations and exemptions seem relevant and well-researched | very clear to read. well aligned with regulations and exemptions seem relevant and well-researched | considering the environmental impact of silicone production it seems aligned to update the criteria with silicone testing | great evolvement of including more packaging criteria. Very well aligned with Nordic Packaging Criteria for Health care products. | Very well aligned with Nordic Packaging Criteria for Health care products and pioneering requirements ahead of the upcoming PPWR | highly relevant to include in criteria | requirement is valid to include and should be part an integral part of a medical device manufacturer's processes | requirement is valid to include and should be part an integral part of a medical device manufacturer's processes | Very clear and easy to follow. Give good guidelines how to respond to the requirements and in what format
| Fresenius Medical Care AG | Germany |  | Many thanks for giving us the opportunity to provide input for the current review process. We support the goal of Nordic Ecolabel criteria towards making a positive environmental impact. We are proud to be a leading company in environmentally advanced products in PD as the first and only continuous license holder since the beginning of this product group. An ecolabel can reward and increase visibility of advanced environmental performers in the product segment. We are committed to delivering dialysis products and services in an environmentally sustainable way. Thus, we support Nordic Ecolabeling`s goal to have ecolabel criteria, which are stricter than EU regulations, highlighting license holder`s commitment to manufacturing environmentally sustainable products. We believe that the existing criteria are still a valid, challenging differentiator within a product segment. There are still only few license holders for this product scope. Which clearly indicates that production of such high-quality medical products in compliance with the existing Nordic Ecolabelling requirements (e.g. without PVC and without SVHC) is still not the standard in the healthcare industry. Fresenius Medical Care products provide life-saving or life-sustaining medical treatment of patients who are directly dependent on the availability and safety of the products described. Our products (medical devices and medicinal solutions) are subject to stringent patient safety requirements. When a substance or a component needs to be substituted by another that has a smaller impact on the environment, the substitute must not compromise patient health and safety. This means the substitute should perform as good as or better than the former raw material/component. The performance and safety of the product and possible alternative components or raw materials must follow strict legal requirements which may include testing, prototyping, clinical trials and validation. This may take several years without the guarantee that it is a viable alternative to begin with. Therefore, medical technology manufacturers first need to have a feasible technical alternative (if at all available) to be able to begin with testing for compatibility, safety, performance, etc. This also includes intensive collaboration with the upstream suppliers. Any new or changed ecolabel criteria can result in conflicts with applicable product safety requirements or time and resource issues for design changes and documentation. Thus, duly balanced ecolabel criteria are needed that ensure advanced eco-design while remaining feasible from the resulting administrative effort and resource perspective. Only this would ensure that ecolabel can be a valuable differentiator for environmental frontrunners. For more detailed, product specific feedback please refer to our additional input by email. | The new proposed name of the criteria “Nordic Ecolabelling for Medical devices in plastic or silicone” is misleading since many of FME Nordic Swan ecolabeled products are in fact medicinal solutions, i.e. not “medical devices”. Medicinal solutions have been a major product group for the establishment of these ecolabel criteria and there should be no doubt about them being in scope also in the future. For consistency we suggest rewording: “Nordic Ecolabeling for plastic or silicone products in medical technology”.
| Region Stocholm_Nordic healthcare actors (NIELS working group) | Sweden, Denmark, Norway, Finland Joint Nordic response from healthcare actors (NIELS working group) |  | We are, as a whole, very satisfied with the criteria document. It is thoroughly developed and highly ambitious. Note: Please also see separate tab with additional consultation responses. | It is stated in the product group definition that products that are mainly, at a minimum of 90% weight, made of plastic or silicone, but later there are criteria for silicone products that are at least 80% by weight. Perhaps there can be an increased clarification in the section for “what can carry the Nordic Swan ecolabel” in the product definition. Closed systems for medical administration e.g. for antibiotics or chemo therapy. Masks: Can you include Laryngeal mask (LMA)? We would prefer you to include bottle teats. Comment to the following: "relevant medical devices in plastic or silicone may be included in the product group upon request".There are large amounts of transparent plastics products (e.g. polycarbonate) used in laboratories today which could instead be produced without fosgen to avoid emissions of chlorides . Have you considered including fosgen free polycarbonate products in the requirements? Can laboratory plastic products also be considered within the scope? Alternatively it may be beneficial to consider for a separate product group. Requirements O11 and O12 are also relevant for other medical products such as electrodes, diabetes products such as glucose monitoring sensor or adhesives for stomy. All adhesives in the medical field should comply with these criteria. Medicine cups with volume markings are Class Im medical devices that store and transport fluid and medicine: Why do you say low RPS? We think these should be included anyways as there is a high consumption and high representation in all kinds of applications. Ophthalmology: Scleral plugs and punctal plugs should also be included Ear, Nose, Throat: Septal plugs | We would wish for more strict requirements for documentation, e.g. that SDS is valid, European and not older than 2022. Is there requirements on whom in the organisation may fill out and approve a), c), d) or e)? Our experience shows that many suppliers/manufactures can't supply a SDS or data sheet. | We think these are great requirements and very ambitious, but we think the market will have difficulties complying. | We think these are great requirements and very ambitious, but we think the market will have difficulties complying. | Great you aim to avoid latex, we think your environmental argumentation is on point. Another point to consider is that there are organic certified Hevea brasiliensis, and natural latex is a good alternative to PVC, as you are also mentioning. Significant ongoing research and industrial development aim to produce hypoallergenic medical products from Hevea brasiliensis. While traditionally associated with Type I latex allergies, modern efforts focus on removing or modifying the allergenic proteins that trigger these reactions. Suppliers to the Nordic market are developing such products, it would therefore be great if certified latex products which do not cause allergies could be certified. As a minimum, certified organic latex should be allowed in products that do not come into contact with skin, until it can be scientifically proven that it does not cause allergies. | Great requirements and positive to focus on cyclical silicone and chloral ones separately in O6. Regarding siloxanes: there are only requirements regarding cyclic siloxanes. Linear siloxanes can also be problematic, like Octamethyltrisiloxane (L3) which in on the candidate list from January 2025. Have you considered also including linear siloxanes in the requirements? Source for more info: https://echa.europa.eu/da/-/echa-adds-five-hazardous-chemicals-to-the-candidate-list-and-updates-one-entry#msdynttrid=rcs5kziY5-LXqh-ebpzlVZsFOFPt-ib-XQxnMUUDaiI |  |  | We would recommend this part to be more detailed: what do you want included in the CO2 emissions? What about raw materials and emission factors for these? What about packaging? Or transportation? Waste treatment? Why one factor for grid electricity? We would e.g. prefer production in Northern Europe to China and could therefore reward production in EU? Maybe you could rephrase the part about electricity grid factor and give more detailed guidelines on which factor to use? E.g. to use 210 gCO2e/kWh for Europe, but use "xx factor" in China etc. with a clear reference on where at European Environment Agency to find that factor. We would recommend you provide a template/appendix to fill out for the manufacturer to ensure correct calculations. E.g. do all manufactures have a detailed enough energy monitoring system? Or do they know how to calculate energy consumption per kg silicone? | Very positive that you include EUH and not only H, we think it is a great list. We are concerned though for substances without harmonised classification. How will you control/approve these? |  | Yes, awesome criteria! We are very positive over O11 and are looking forward to using these requirements as these are very important for patient safety. | Very positive criteria as well. | Good criteria as well! Great that you include quaternary ammonium compounds, which are not readily aerobic biodegradable. We would like requirements to avoid acrylates here. Are these included in any of your list already? |  | We really like your references to you position statements. Don't you have requirements to minimise the amount of packaging materials used? Weight of packaging vs. weight of product? Is it possible to include filling materials? Regarding FSC/PEFC documentation: We would recommend to include CoC number and certificate and traceability on data sheet/invoice or similar. | We really like the reference to marked pictograms. | We would recommend to also include the requirement for CE and notified body in criteria document for Protective and Absorbent Hygiene Products. |  |  | As mentioned under O8 we would recommend a template for calculations of CO2 emissions.
| Coloplast A/S | Denmark |  | Some of the requirements – for example a total ban on chemicals - may offer only limited environmental impact relative to the disproportionate documentation burden they would demand from companies and raw material suppliers/manufactures. The medical device industry is highly regulated, with strict requirements under MDR and REACH, ensuring patient safety and environmental responsibility. This approach supports sustainability and innovation without imposing unnecessary complexity on an already tightly controlled sector. Suggestion for additional requirements for certification of manufacturing site: • ISO 13485 (Medical Devices - Quality management systems – Requirements for regulatory purposes) to strengthen the following requirements: • O17: Safety • O18: Customer complaints • O19: Traceability • ISO 14001 (Environmental management system) for: • Systematic approach for managing and reducing the environmental impact from manufacturing of products that are labelled with the Nordic Swan Ecolabel. • ISO 45001 (Occupational health & safety management system) for: • Systematic approach for managing and improving the occupational health and safety for the employees manufacturing products that are labelled with the Nordic Swan Ecolabel. | No comments | No comments | No comments | No comments | No comments | We support the 100-ppm limit for each of D4 (CAS 556‑67‑2), D5 (CAS 541‑02‑6) and D6 (CAS 540‑97‑6) and agree that verification testing is essential. To ensure feasible, auditable implementation across elastomers and silicone fluids/oils (including silicone used in adhesives or surface treatments), we propose Nordic Swan Ecolabel to accept method flexibility: The CES–Silicones Europe methods or any ISO/IEC 17025‑validated GC‑MS method achieving LOQs at or below the target level. | No comments | No comments | No comments | The comments apply to O9-O14. Some of the requirements – for example a total ban on chemicals - may offer only limited environmental impact relative to the disproportionate documentation burden they would demand from companies and raw material suppliers. The medical device industry is highly regulated, with strict requirements under MDR and REACH, ensuring patient safety and environmental responsibility. This approach supports sustainability and innovation without imposing unnecessary complexity on an already tightly controlled sector. | No comments | No comments | No comments | No comments | No comments | We would need clarifications and definitions of the following: • Define what “recycled*” refers to. • Provide examples of FSC/PEFC-schemes. • The 70 % corresponds to wood weight in “O15 Packaging materials” and the text in “Background to O15 Packaging materials” is not defined the same way. We suggest that both sections should be defined like “A minimum of 70wt% of the wood raw material in ….”. | With respect to the environmental impact versus the effort, we propose to wait for updating artwork of the packaging cardboard until the EU requirement for pictograms (EU2025/40) is specified. | This applies to O17-O19: We propose to include the following certification in the requirements: • ISO 13485 (Medical Devices - Quality management systems – Requirements for regulatory purposes) to strengthen the following requirements: • O17: Safety • O18: Customer complaints • O19: Traceability | See comments on O17 | See comments on O17 | No comments
| Institutet för Miljömedicin, Karolinska Institutet | Sweden |  | Institutet för Miljömedicin (IMM) är positiva till de nya förslagen som generellt stärker kriterierna för särskilt problematiska kemikalier, och framförhållningen att kriterierna även ska täcka problematiska kemikalier som inte finns i produkter på marknaden idag. Vi har specifikt granskat avsnitt 5.4 ”Additives in polymer materials” (O9-O10) och 5.5 ”Adhesive and surface treatment” (O11-O14). | [9 empty] | Det är otydligt varför inte att samma faroklasser är listade som Excluded hazards i O9, O11 och O12 (Tabell 1, 2 och 3). Till exempel, • ’hazardous to the aquatic environment’ och ‘acute toxicity’ listas som Excluded hazards i O11 men inte i O9 eller O12 • ’STOT SE 1’, ’STOT SE 2’, ’STOT RE 1’ och ’STOT RE 2’ listas för O9 och 11 men för O12 listas endast ’STOT RE 1’. Anledningen till att olika faroklasser inte är relevanta för vissa kriterier kan förtydligas i texten. | Vi stödjer kravet att potentiella och identifierade hormonstörande ämnen listade i “Endocrine Disruptor Lists" List I, II eller III inte får ingå i produkterna. Vi föreslår att förtydliga att kravet gäller ämnen identifierade antingen som hormonstörare för ENV eller HH, eller både och. Vi föreslår också att lägga in en referens/hänvisning som förtydligar vilka listor det är som avses, t ex till websidan: https://edlists.org/the-ed-lists Background to O10: Här listas olika kemikalier/kemikaliegrupper som inte får ingå i produkterna. Texten kan ses över och göras tydligare och mer enhetlig för de olika ämnena. Till exempel: • För ftalater: det är oklart vad som menas med ”…there is a ban on this group of substances”. Det kan förtydligas om det är ett förbud i Svanenmärkta medicintekniska produkter i plast eller silikon som avses, specifikt. • För bisfenoler: det är oklart vad som menas med ”… are now prohibited.” Det kan förtydligas om det är ett förbud i Svanenmärkta medicintekniska produkter i plast eller silikon som avses, specifikt. Vi föreslår att Quaternary ammonium compounds och Volatile aromatic compounds även listas för O10 (nu endast listade under O13). | Det är otydligt varför inte att samma faroklasser är listade som Excluded hazards i O9, O11 och O12 (Tabell 1, 2 och 3). Till exempel, • ’hazardous to the aquatic environment’ och ‘acute toxicity’ listas som Excluded hazards i O11 men inte i O9 eller O12 • ’STOT SE 1’, ’STOT SE 2’, ’STOT RE 1’ och ’STOT RE 2’ listas för O9 och 11 men för O12 listas endast ’STOT RE 1’. Anledningen till att olika faroklasser inte är relevanta för vissa kriterier kan förtydligas i texten. | Det är otydligt varför inte att samma faroklasser är listade som Excluded hazards i O9, O11 och O12 (Tabell 1, 2 och 3). Till exempel, • ’hazardous to the aquatic environment’ och ‘acute toxicity’ listas som Excluded hazards i O11 men inte i O9 eller O12 • ’STOT SE 1’, ’STOT SE 2’, ’STOT RE 1’ och ’STOT RE 2’ listas för O9 och 11 men för O12 listas endast ’STOT RE 1’. Anledningen till att olika faroklasser inte är relevanta för vissa kriterier kan förtydligas i texten. | Vi stödjer kravet att potentiella och identifierade hormonstörande ämnen listade i “Endocrine Disruptor Lists" List I, II eller III inte får ingå i produkterna. Vi föreslår att förtydliga att kravet gäller ämnen identifierade antingen som hormonstörare för ENV eller HH, eller både och. Vi föreslår också att lägga in en referens/hänvisning som förtydligar vilka listor det är som avses, t ex till websidan: https://edlists.org/the-ed-lists Background to O13: Här listas olika kemikalier/kemikaliegrupper som inte får ingå i produkterna. Texten kan ses över och göras tydligare och mer enhetlig för de olika ämnena. Till exempel: • För ftalater: det är oklart vad som menas med ”…there is a ban on this group of substances”. Det kan förtydligas om det är ett förbud i Svanenmärkta medicintekniska produkter i plast eller silikon som avses, specifikt. • För bisfenoler: det är oklart vad som menas med ”… are now prohibited.” Det kan förtydligas om det är ett förbud i Svanenmärkta medicintekniska produkter i plast eller silikon som avses, specifikt.
| Miljøstyrelsen | Denmark |  |  |  |  | Halogeneret plast. Ingen kommentarer. | Halogeneret butyl gummi. Ingen kommentarer. | Naturlig gummilatex. Ingen kommentarer. | Silikone. Ingen kommentarer. | Produktion af silikone. Emission af silikonestøv og klorider til luft. Emission af støv (side 14): 1a) Støvreducerende tiltag til luft (BAT-AEL), eller 1b) Overholde grænseværdi på 5 mg/Nm3 (årligt gennemsnit) 2a) Kloridreducerende tiltag (side 15): Røggasser skal undergå termisk oxidering for at reducereudslip af klorider. Miljøstyrelsen bemærker, at emissionskravene er baseret på EU Ecolabel miljømærkekriterier, og har ikke yderligere kommentarer. | Produktion af silikone. Emission af kobber og zink til vand. Vandudløbet fra produktionstrinnet for polydimethylsiloxan (PDMS) skal forbehandles ved udfældning eller flokkulering under alkaliske forhold, efterfulgt af sedimentation og filtrering og følgende emissionskrav: Max-konc. af kobber i udløb fra forbehandlet spildevand: < 0,5 mg/l. Max-konc. af zink i udløb fra forbehandlet spildevand: < 2 mg/l. Miljøstyrelsen er enig i, at det er hensigtsmæssigt at reducere indholdet af tungmetaller, herunder kobber og zink, i spildevandet inden afledning til renseanlæg. Miljøstyrelsen bemærker, at emissionskravene er baseret på EU Ecolabel miljømærkekriterier for menstruationskopper af silikone, og på baggrund af BREF/BAT for silikoneproduktion (BREF 2007). Såfremt produktionen af silikoneråmaterialet skulle foregå i Danmark, vurderer Miljøstyrelsen, at de foreslåede max-koncentrationer ikke ville være tilstrækkelige til at varetage miljøhensyn iht. vandrammedirektivet ved afledning til renseanlæg eller udledning til recipient. | Emission af CO2. Ingen bemærkninger. | Klassificering af indgående stoffer i additiver. Miljøstyrelsen er enig med Nordisk Miljømærkning, at det er uhensigtsmæssigt at tillade antioxidanter, der er klassificeret som reproduktionstoksisk kat. 2 i svanemærkede produkter, på trods af at de tilsættes for at undgå afgasning af N-nitrosaminer. I baggrundsinformationen fremgår det ikke hvorvidt, der findes substitutionsmuligheder til disse antioxidanter. Miljøstyrelsen så gerne, at svanemærkets krav i højere grad skabte et incitament for udvikling og anvendelse af antioxidanter, der ikke har denne eller lignende skadelige effekter. Derudover bør baggrundsafsnittet uddybe, hvornår antioxdanter er nødvendige, herunder om det er lovgivning, samt hvornår de ikke er, og om det altid bør være REP. Cat 2 klassificerede antioxidanter. | Udelukkede stoffer i additiver. Kriterierne udelukker PFAS-stoffer og halogenerede organiske kemikalier. Imidlertid er de skrevet ind hver for sig, selvom PFAS tilhører gruppen af halogenerede organiske kemikalier. Miljøstyrelsen mener, at det er fagligt set mere korrekt, hvis PFAS tilføjes som en del af en overordnede gruppe af halogenerede organiske kemikalier under kriterium Udelukkede stoffer, fx: Halogenated organic compounds • Per- and polyfluoroalkyl substances (PFAS) • Organic chlorine compounds (e.g., short-/medium/-long chained chloroparaffins) • Other halogenated organic compounds* Det er uklart, hvorfor der er behov for undtagelsen for “Pigments that meet the EU’s requirement concerning colourants in food packaging under Resolution AP (89) point 2.5”. Those pigments contain small amounts of PCBs, which in general should be avoided.” Som bemærket under kriterie O9 bør nødvendigheden af en undtagelse uddybes (med evt. henvisning til lovgivningskrav), og herunder beskrive mulige alternativer. I forhold til undtagelsen for halogenerede organiske pigmenter formoder Miljøstyrelsen, at denne undtagelse ikke gælder PFAS-baserede pigmentstoffer. | Klassificering af klæbemidler og blandinger til overfladebehandling. UV-hærdende akrylater er undtaget fra dette krav på betingelse af at applikationen sker i et lukket system. Akrylater kan være hudsensibiliserende indtil de er fuldt hærdede. Derfor anbefaler Miljøstyrelsen, at kriteriet tilføjes, at akrylater kun kan undtages dette krav, hvis det kan påvises, at akrylaterne er fuldt hærdede efter applikation i et lukket system og før arbejdstagere kommer i kontakt med klæbemidlerne/overfladebehandlingen. | Klassificering af indgående stoffer i hhv. klæbemidler og blandinger til overfladebehandling. Ingen kommentarer. | Udelukkede stoffer i hhv. klæbemidler og blandinger til overfladebehandling. Ingen kommentarer. | Silikone i hhv. klæbemidler og blandinger til overfladebehandling. Ingen kommentarer. | Emballage og klorbaseret plastik. Miljøstyrelsen bemærker, at der ikke er sat kriterier for plastmaterialet, hvis man ser bort fra at PVC og PVDC ikke må anvendes. Kriterierne synes lempelige i forhold til at sikre genanvendelse og bør fokusere på, at emballagen til svanemærket medicinsk udstyr i så vidt muligt omfang er designet til genanvendelse; medmindre der er sundhedsforhold, der taler imod en højere grad af genanvendelse. Vedrørende definitioner bør de anførte definitioner (for primær, sekundær og tertiær emballage) justeres til at passe med Emballageforordningens definitioner: »salgsemballage«: emballage udformet på en sådan måde, at den udgør en salgsenhed bestående af produkter og emballage til slutbrugeren på salgsstedet. »multipakemballage«: emballage udformet på en sådan måde, at den udgør en samling af et vist antal salgsenheder på salgsstedet, uanset om denne samling af salgsenheder sælges som sådan til slutbrugeren, eller om den tjener til at lette opfyldning af hylder på salgsstedet eller til at skabe en lager- eller distributionsenhed, og som kan fjernes fra produktet, uden at dette ændrer produktets egenskaber. »transportemballage«: emballage, der er udformet på en sådan måde, at håndtering og transport af en eller flere salgsenheder eller en samling af salgsenheder lettes for at forhindre skade på produktet forårsaget af håndtering og transport, men som ikke omfatter vej-, jernbane-, skibs- og luftfragtcontainere. | Designkrav til emballage med henblik på genanvendelse. Det er uklart hvordan adskillelse af komponenter skal efterprøves og dokumenteres. Der er muligvis behov for at aligne krav med de udvidede producentansvarsregler, som er indført
| Wellspect | Sweden |  | In MedTech product development, approximately two years are required for activities such as testing, biocompatibility assessments, and aging studies. This timeline would become necessary if an existing material were to be prohibited and a replacement material needed to be developed. | [9 empty] | It is usually difficult to get information on additives used in plastic materials. So this will be tricky. It could be that additives are added at multiple positions along the chains. | A lot of different lists to check against for those who should sign any statement (Appendix 4). It will be necessary to have relevant links easily accessible at Svanen, preferably arranged by requirement (O-number). Even if the lists are detailed there will be instances where it will be difficult to know if certain substances are excluded or not. |  |  | A lot of different lists to check against for those who should sign any statement (Appendix 4). It will be necessary to have relevant links easily accessible at Svanen, preferably arranged by requirement (O-number). Even if the lists are detailed there will be instances where it will be difficult to know if certain substances are excluded or not.
| Swedish Medtech, Sailab, Melanor, Medicoindustrien | Sweden, Norway, Finland, Denmark |  | On behalf of the Nordic medical technology industry associations, Swedish Medtech (Sweden), Melanor (Norway), Medicoindustrien (Denmark) and Sailab (Finland), I would like to submit our joint response to the public consultation “Medical devices in plastic or silicone (098)”. The reply will be sent through e-mail as separate pdf. Note: Please see seperate tab with consultation responses. | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
| VinylPlus Healthcare, PVC Forum Sverige and PVC-Informationsrådet | Europe, Sweden, Denmark |  | Note: Please see seperate tab with consultation responses. |  |  | Note: Please see seperate tab with consultation responses.

## Nordic healthcare actors (NIELS
Joint Nordic response from healthcare actors (NIELS working group) | Answer from Nordic Ecolabelling:
1. In short: set criteria for energy mapping and/or set a minimum % of renewable energy requirement at the manufacturing stage. Longer comment: While the RPS assumption regarding energy consumption being higher than in raw material phase compared to manufacturing phase may be reasonable to make, there still appears to be a lack of data and certainty (as is made clear in the assessment knowledge about energy consumption in manufacturing is low). This could provide to be a valuable gap to cover by including a criteria on energy mapping, either exactly or similar to the energy mapping criteria in the criteria group for floor coverings. That way the Nordic Swan may gain further insight and facts on energy use for future revision of this product group. The energy mapping would then require an energy audit according to standards, and energy mapping, to be conducted or verified by a third party, as well as an action plan with purpose to reduce energy consumption. In addition to the above it would be suggested to identify the energy mix used for manufacturing (fossil vs. renewable), and that an action plan could also aim at transitioning to renewables. The benefits of going with this approach overcomes the issue of having to set any specific limit for energy production per product, as this covers a wide range of different products. | A new requirement for energy mapping and energy action plan at the manufacturing site of the final product has been added.
2. Have you considered to take into account the location of production? You conduct a site audit, so it would be great to set requirements to local energy source and renewable energy consumption (see also above) that is used in production (you only partly have that covered under O8, specifically for silicone) or to local waste handling and waste water treatment. | We have decided to focus on energy mapping and energy action plan, see above.
3. We think it is great you aim for 100% PVC free products. We have some general questions and comments though: We can tell you aim for silicone as alternative to PVC. But as you stated yourselves, the market does not have alternatives for PVC in all product groups. What alternatives do you expect for e.g. tubes? And what are the risks in these alternatives? Have you e.g. considered to restrict use of harmful isocyanates in PU tubes (we think most are included in harmonised CLP, but not sure)? Or use of styrene or trace metals in catalysts in TPE products? | Only poduct types that have shown medium to high RPS are included in the criteria. For all polymer materials there are strict requirements for additives. Other than D4, D5 and D6 in silicon, we have not set requirements for other impurities that are not comming for additives.
4. The Nordic healthcare regions request for the Nordic Swan to be able to provide concrete information on the environmental/climate benefits of choosing a Typ 1 Ecolabelled product vs. Choosing a non-ecolabelled product (lets say conventional). This would be by clearly providing information on these comparisons or/and providing public procuring authorities with a general calculation method that we may apply. This is important because going further in criteria setting in public procurement requires motivation and argumentation for why we should go for this (for internal acceptance and buy-in, as well as market understanding), and the best way to provide a good argument is show clear benefits on paper. We need to easily detect and determine what we gain from buying the ecolabelled products instead of non-labeled. | Nordic Ecolabelling can provide information about environmental argumentation for Nordic Swan Ecolabelled medical devices. It will not be possible to provide a general number for the environmental/climate effect by choosing ecolabelled medical devices genereral, because the criteria cover many different products types and materials, and lack of market benchmarkers to compare with. It will not be possible to provide a general calculation method because of e.g. lack of benchmarking and chemical requirement can usually not be calculated.
5. Page 49 Mecho Scheme on materials - The following is assessed on End-of-life in the MECHO scheme on materials: "Loss of resources: Medical disposable products are mainly sent for incineration due to contaminated waste. Material recycling is at present non-existent but may be possible in the future". This is not exactly correct, while this is very much the general/major flow, material recycling is already possible today but occurs to a very limited extent. There are some circular examples of circular protective aprons in Region Stockholm done at large volume, made out of PE. Secondly there are technical solutions already today for decontaminating a lot of the hazardous/contaminated waste, making it possible to enter material recycling streams (there are still some hurdles to overcome in terms of the handling of the waste streams at hospitals as well as willingness of reycling actors to receive this sort of product groups for recyling). Thirdly, there are also cases of successful separation collection and recycling of medical PVC, instead of going to incineration (see Recovinyl, and the following link https://url41.mailanyone.net/scanner?m=1voMc3-00000000794-3ujo&d=4%7Cmail%2F90%2F1770387000%2F1voMc3-00000000794-3ujo%7Cin41t%7C57e1b682%7C14161982%7C11838427%7C6985F76F28C6D89E0A1264D99CA6AB1C&o=%2Fphti%3A%2Fvtshpnyausell.cltureehans%2Fsbaiautitily%2Fvniyeulp%2F-mdls&s=Doybli1xj0i7RBf0x-jJJgZWPZk) . While not at large scale solution in practice today, it is technically possible, and a lot of efforts at making it happen are ongoing, this should be reflected in the MECO scheme analysis for End-of-life stage. Fourthly, we are also seeing the development of take-back schemes on various types of medical plastic products from suppliers, which has been observed from our countries. Finally, we would like the Nordic Swan to be up-to-date about potential changes in the MDR regulations that may also have a bearing on this criteria group, as it is important to take account and be aware of regulative changes. Please see Proposal for a regulation to simplify rules on medical and in vitro diagnostic devices published 16 December 2025. https://health.ec.europa.eu/publications/proposal-regulation-simplify-rules-medical-and-vitro-diagnostic-devices_en | We have changed the text in the MECO to "Material recycling is at present low but may be expanded in the future as several initiatives work on this. There is also development of take-back schemes, that may be further developed in the future". At this point it is too early to consider to set requirements in this area, but we will follow the development.
6. page 12 on following text: Background to O2 Halogenated plastics "Halogenated plastics such as PVC are not allowed to reduce problems in waste handling and to reduce the risk of health-related problems". Firstly check the language as it seems the syntax is a bit wrong and misleading (an overall english language check is recommended for the final version). Secondly, while agreed that it is not a good outcome for PVC products to go to municipal waste incineration, there seem to be accounts of modern waste incineration facilities stating they have a certain (limited) capacity for burning PVC. This may be worthwhile researching and revising according to findings. | Nordic Ecolabelling has updated the background documentation text and references for the argumentation for not accepting PVC as a material in medical devices. The requirement itself has not been updated.

## Swedish Medtech, Sailab, Melano
Response from Swedish Medtech, Sailab, Melanor, Medicoindustrien | Answer from Nordic Ecolabelling:
Please see answer in tab "Responses and feedback".

## VinylPlus Healthcare, PVC Forum
Response from VinylPlus Healthcare, PVC Forum Sverige and PVC-Informationsrådet | Answer from Nordic Ecolabelling:
Please see answer in tab "Responses and feedback".